## Purchased Parts
Sr.No. | Part Name | Description | Quantity | Comments/Links | Where to Purchase?
[5 empty] | Vendor 1 |  |  | Vendor 2
[5 empty] | Address & Contact Details | Total Price(INR)
[6 empty] | MRP | Bulk | Address & Contact Details | Total Price(INR)
1 | O2 Inlet Tubing + Adjustable Venturi Valve (1137015 or 1117015) | For Oxygen Inlet and Air Mixing | 1 | Multi Ox Venturi Mask kit(1117015) is preferred but we have used 1137015 Port - 22F | A V Surgitech, Shop.No.23, Plot No-272 - 273, Hind Rajasthan Shopping Bhaudaji Road, Sion West, Sion , Mumbai - 400022 Kishor - 9137920335 022-24022999 | 260 | 130
2 | Viral Filter (1344007S) | Filtering inlet and outlet air to the patient | 1 | Ports: Tube End 22M, Mask end 22F | A V Surgitech | 375 | 182
3 | Y connector (1926000) | As a coupler to connect inlet and outlet tubes on the single port mask | 1 | All 3 ports are 22M | A V Surgitech | 360 | 250
4 | FlexTube Ventilator Tubing(2000000) - (Y connector + 1.6m x 2) | Tubing to connect parts | 1 | Tubing is 22F, Y connector is 22M | A V Surgitech | 795 | 250
5 | Single Port Anaesthetic Face Mask (1516000) | Single port Mask | 1 | Ports: 22M - Inlet | A V Surgitech | 1025 | 500
6 | CPAP Mask Harness (2224000) | To fit the mask tightly to the patient | 1 |  | A V Surgitech | 1200 | 560
7 | Adjustable PEEP Valve (2226000) | Adjustable Pressure valve to remove air during the exhaust cycle | 1 | Port: 22M | A V Surgitech | 2000 | 800
8 | Rebreathing Bag 2L (2820000 ) | Reservoir Bag as buffer | 1 | Port: 22F Neck | A V Surgitech | 675 | 380
9 | Intersob Plus Sodalime 5L Can (2180000) Pink to White or (2179000) White to Violet | Sodalime for Co2 Removal from the system | 1 |  | A V Surgitech | 3200 | 1450
10 | Plastic Jar 2L | Container for holding sodalime | 1 | On Outside: 4 Ports (3 ports with directional valves) - All 22M On Inside: 2 Ports: Both 22M to connect tubing | Amazon | 100 | 100
11 | Bolts | Fastening the 3D printed ports to the lid cover and diffuser | 39 | M3 x 40mm x 18pcs M3 x 16mm x 6pcs M3 x 12mm x 15pcs | Local Hardware shop | 75 | 75
12 | Nuts |  | 39 | M3 x 39pcs | Local Hardware shop
13 | Fine mesh Plastic Net (Approx 90mm diameter) | To spread the incoming air throughout the width of the can. To prevent caking of sodalime | 2 | Repurpose from a fine mesh mosquito net or plastic tea strainer | Local Hardware shop | 30 | 30
14 | Circular silicone disc | For making unidirectional valves | 3 | Diameter - 26mm. Currently repurposed from N95 masks with valves | Local Medicalshop for N95 masks. Need to find a proper vendor who makes silicone discs | 75 | 75
[5 empty] | Total (INR) | 10170 | 4782

## 3D PrintLaser Cut Parts
3D Printed Parts
Sl No | Part Name | Description | Quantity | Material | Approx 3D Print Time(in mins) Infill 100% | Comments/Links
1 | Diffuser_Bottom_Cover | To close the plastic mesh assembly | 1 | ABS or PLA | 200 mins
2 | Diffuser_Filter_Holder | To hold the plastic mesh on to the ring | 1 | ABS or PLA | 125 mins
3 | Diffuser_Top_Cover | To hold plastic mesh in the assembly | 1 | ABS or PLA | 180 mins
4 | Male_Adapter_Port_22M | Ports on the jar for various connections | 6 | ABS or PLA | 120 mins each
Laser Cut Acrylic Parts
Sl No | Part Name | Description | Quantity | Material | Thickness | Comments/Links
1 | Port_Spacer | Used as spacers in the valve design | 4 | Acrylic | 5 mm
2 | Valve_Spacer | To be used in betwen Valve_Base_Plate and Valve_Top_Plate parts | 6 | Acrylic | 5 mm
3 | Valve_Base_Plate | Valve sits on this plate | 3 | Acrylic | 2 mm
4 | Valve_Top_Plate | It keeps valve in place using nut and a bolt | 3 | Acrylic | 2 mm
5 | Net_Filter_Holder | To hold the Net filter kept at bottom on container lid | 3 | Acrylic | 2 mm